## Slide 1: An Introduction to the Atomic Force Microscopy (AFM)
National Technical University of Athens
Athens 2022

## Slide 2
About Microscopy
Derived from the Greek words  μικρός + σκοπεīν
Zacharias Janssen (1585 – 1632)  the first fully functional microscope
1931 Ernst Ruska began constructing the earliest electron microscope, Transmission Electron Microscope (TEM)
1936 Erwin Müller constructed the first Field-emission Microscope (FEM)
1937 Manfred von Ardenne constructed the first functional Scanning Electron Microscope (SEM)
1981 Gerd Binnig & Heinrich Rohrer developed the Scanning Tunneling microscope (STM).
1986 Gerd Binnig & Quate & Gerber invented the Atomic Force Microscope (AFM)

## Slide 3
AFM Introduction
Allows the study and observation of the structure of surfaces with excellent precision.
It differs from the other microscopy methods, as it is not based on creating images from a focused energy beam but on the natural contact with the surface.
It does not produce 2D but 3D images.
It can measure almost every type of specimens without special treatment (metals – ceramics – polymers – cells – DNA and RNA molecules...).

## Slide 4
AFM Introduction

## Slide 5
Simple idea  difficulty in construction due to the requirement of ultra high precision
Basic parts:
  Microscope stage
    Scanner  the probe is moving relatively to the specimen
    Sample holder
    Force sensor
  Control electronics
  PC
Optical microscope
Vibration isolation platform
AFM Principles of operation

## Slide 6
Piezoelectric transducers
  Electromechanical devices, which convert electric potential to movement and vice versa (0.1nm/V).
  They are used for the precise control of probe displacement.
Force transducer
  It measures the force occurring due to the contact between the probe and the specimen surface.
  It is composed of various devices.
  Typical force values: 10 piconewton
Feedback control
  It is used to maintain constant force between the probe tip and the specimen surface.
AFM Principles of operation

## Slide 7
AFM Principles of operation
Feedback control
Force transducer
Piezoelectric transducer

## Slide 8
Extremely sharp probe tip  ultra high resolution
Capability of macroscopic displacement of the probe.
Detection of force in the order of nN (or even smaller) in order to avoid breakage of the probe.
Appropriate feedback control allowing rapid control of the probe.
Piezoelectric transducer for scanning in XYZ axis.
Considerably robust structure.
AFM Challenges of the construction

## Slide 9
AFM Stage
Sample-scanning (left) and probe-scanning (right) microscopes

## Slide 10
AFM Stage – piezoelectric scanners
Typical geometries of piezoelectric elements.
(PdBaTiO3 - Pb[ZrxTi1–x]O3)
Four external electrodes control the movement in x-y axes and an internal electrode controls the z axis (tube scanner)
Flexure scanner movement in xy axes
The simplest motion device for the three axes.

## Slide 11
AFM Stage – force sensors
Scanning Tunneling Microscope (STM)
Michelson interferometer
Crystal oscillator
Piezoresistive cantilevers
Optical lever sensor

## Slide 12
The probe is fabricated by “MEMS Technology” (it is considered disposable)
The quality of image depends on the geometry and characteristics of the probe  no residual stress should occur, the probe should not be brittle and should be sufficiently sharp (10 nm)
Probe materials: silicon (Si) & silicon nitride (Si3N4)
AFM Cantilevers & Probes
SEM images of different types of sharpened silicon probes. A: a standard silicon oscillating mode probe. B: ‘super-sharp silicon’ probe sharpened with an electrochemical etch. C: ‘high-aspect-ratio’ probe sharpened with ion milling. D: probe modified with carbon nanotube

## Slide 13
AFM Cantilevers & Probes
Contact and non-contact probes. Left: a typical v-shaped contact-mode cantilever. The whole probe is made from silicon nitride (Si3N4), and has an integrated square pyramidal probe tip. Right: a probe designed for oscillating modes such as non-contact AFM. The cantilever is usually rectangular (or a modified rectangle shape like this one). The whole probe is made from silicon, and is much stiffer and more prone to breaking than the contact probe; however it has a sharper tip.
Cantilever calibration is necessary
Various calibration method
  Sader  geometric characteristics
  Cleveland measurement of the alteration of structure eigenvalues after mass addition
  Thermal noise

## Slide 14
Topographic Modes
  Contact mode
  Oscillating modes
    Non-contact / close contact
    Intermittent – contact mode
Non-topographic Modes
  Force spectroscopy
  Nanoindentation
  Mechanical property imaging
    Lateral force microscopy
    Phase imaging
  Magnetic force microscopy
  Electric force microscopy and scanning Kelvin probe microscopy
  Electrochemical AFM
  Thermal modes
  Surface modification
AFM Modes

## Slide 15
The simplest mode of operation, which is the “basis” for other operation modes.
Capabilities of ultra high resolution in reasonable time.
It is based on the potential which occurs in “very small distances”.
  Zero force is exerted away from the surface  zero deformation
  As the probe approaches the surface, attractive forces are exerted  existence of deformation As the probe approaches even closer, repulsive forces are exerted  existence of deformation
The goal is the interaction between probe and surface to be in the region of repulsive forces.
The pressure which is exerted and the displacement of the cantilever depends on the feedback system and the deviation from a set point.
AFM Contact Mode

## Slide 16
Two signals are recorded
Α) The deviation of the signal from the set-point signal (error signal).
B) The motion of the piezoelectric mechanism so that the force between the probe and the surface remains constant (constant force mode).
If feedback is turned off while imaging, the AFM operates in constant height contact mode instead of constant force.
AFM Contact Mode
Illustration of the relation between height and deflection images. Left to right: height, right-shaded height and deflection images of the surface of a mosquito eye. The height image shows how much the z scanner moves to maintain the set-point. The deflection image shows how the cantilever bends as it passes over the sample, and is the signal used for feedback in contact-mode AFM. This image is very similar to the shaded height.

## Slide 17
It is used when high resolution and sharp images are required.
Attention should be paid due to the repulsive forces between the probe and the specimen.
  Damage is caused on the specimen surface  it is not intended for damage susceptible specimens (soft, easily deformed…)
  Wear / damage of the probe
Lateral forces are developed due to contact and movement.
Attention should be paid to high forces being exerted on the probe (>nΝ)
Ideal for microscopy observations in fluids.
AFM Contact Mode

## Slide 18
AFM Contact Mode
The effect of different feedback settings. The motion of the probe over a simple sample, resembling a calibration/test grid (left), and the oscilloscope window showing deflection (z-error) and height information (right). Note that feedback in AFM is never instantaneous, so the bottom example still shows some imperfections

## Slide 19
The cantilever is oscillating (usually at its resonant frequency) by the aid of a piezoelectric device.
The interaction between the probe and the specimen surface is causing damping  reduction of oscillation frequency and amplitude.
The oscillation is detected by a force transducer.
The distance between the probe and the specimen remains constant by an appropriate feedback loop in similar way as in the Contact Mode.
There are different modes with different oscillation amplitudes and different approaches for the prediction of oscillation deviations.
AFM Oscillating Modes

## Slide 20
The choice of oscillation amplitude is related to the existence (or non-existence) of the region of the attractive forces.
  Restricted oscillation amplitude (10 nm)  non-contact mode (A)
  Large oscillation amplitude (<100nm) intermittent mode (B)
Non Contact Mode
High sharpness
Operation in the region of attractive potential (attractive Van der Waals forces)  low interaction forces between probe and specimen surface.
Use of a cantilever with high stiffness (300 – 400 kHz)
Control of dynamic phenomena.
AFM Oscillating Modes

## Slide 21
It can be applied to almost every specimen.
It yields more precise results compared to the Intermittent mode due to lower wear of the probe.
Difficulties due to existence of a “contamination layer”  forces due to capillary phenomena
Oscillation over the “contamination layer”
(+) detection and observation of droplets
(-) lower sharpness
Oscillation inside the “contamination layer”  higher sharpness but more difficult
Use of Ultra High Vacuum (UHV)  atomic resolution (Morita et al.)
AFM Oscillating Modes

## Slide 22
Intermittent Mode
The interaction between probe and specimen surface moves through all regimes (neutral-attractive-repulsive).
Almost zero side forces.
The probe can possibly penetrate the “contamination layer”
Increase need for collecting data from the feedback system.
Due to the existence of repulsive forces, wear can probably occur on the specimen surface or/and the probe.
The use of high order eigenvalues for sensitivity analysis and better feedback.
AFM Oscillating Modes

## Slide 23
Force Spectroscopy
The specimen remains fixed in xy axis and can be displaced only in z-axis in order to measure interatomic forces.
Creation of force-distance graphs.
Capability of appropriate modification of the probe tip according to specific requirements.
Parameters which affect the measurements:
  Number of interacting atoms
  Orientation of atoms
  Approaching speed
  Measurement environment
  High variability of results
AFM Non-topographic Modes

## Slide 24
Nanoindentation
Recording of force-displacement curve.
Advantages
  High sensitivity
  Inherent capability of measurement of imprint Capability of position control with high precision
Disadvantages
  Non-vertical approach of the probe
  Nonlinearity of displacement in the z-axis
  Mandatory calibration for measurement of real forces
AFM Non-topographic Modes

## Slide 25
Mechanical property imaging
Lateral force microscopy: Measurement of “lateral” deviation of the signal  data regarding the surface profile and friction coefficient
Phase imaging: Measurement of the phase shift of the signal in Intermittent – contact mode due to adhesion and viscoelasticity phenomena
AFM Non-topographic Modes
Examples of phase contrast in IC-AFM on different samples. Top: a triblock copolymer topography (left) barely shows height differences for the different phases. The phase image (right) shows clear contrast. Bottom: Langmuir–Blodgett film on mica, the high topography region (the monolayer) has a higher phase contrast than the mica in the phase image. This image shows how the edges of these phases also show different contrast in the phase image, due to changes in tip–sample contact area.

## Slide 26
Magnetic force microscopy
Capability of magnetic properties measurement
Use of a magnetic probe  typical silicon probe with magnetic material coating (Co, NiCo, CrCo)
The magnetic forces are several order of magnitude lower than the interatomic forces measurement from a constant distance  different methods (most of them are measuring both topography and magnetic field in consecutive scans)
AFM Non-topographic Modes
Example MFM images (10 μm x 10 μm). Left: topography of magnetic tape sample. Right: MFM image of the same region, showing magnetic fields above recorded data bits on the tape.
Schematics of various implementations of MFM. A: lifting probe between topography and MFM images. B: Bard method of lifting lever between scan lines. C: z set-point oscillation. D: Hosaka method of moving probe close to surface, and recording MFM signal at various points for each height.

## Slide 27
Electric force microscopy & scanning Kelvin probe microscopy
As with the MFM, electric fields are measured
In SKPM electric potential with a specific frequency (eigenfrequency of the cantilever) is applied on the probe.  The different potential between the probe and the surface causes electrostatic forces which lead to cantilever oscillations.  Feedback loop via application of potential equivalent to that of the surface tends to reduce this difference to zero.
Electrochemical AFM
Alterations in topography due to chemical reaction after the application of electric potential.
The potential is regulated for the oxidation or reduction of the surface.
AFM Non-topographic Modes
Electrochemical AFM example. Images showing the morphology of a CdTe film during electrochemical deposition of Au, at various times (as shown in figure) at a potential of -0.35 V

## Slide 28
Surface modification
Use of AFM for manufacturing – modification of the surface in –nano level
Different techniques
  Localized oxidization
  Engraving
  Dip pen nanolithography
  Thermomechanical writing
Thermal modes
It is used for the measurement of thermal properties of the specimen.
The probe is used either as a resistance for localized heating or as a thermometer.
AFM Non-topographic Modes
Examples of AFM-based lithography. Left: polymeric patterns on silicon formed by anodic oxidation, showing line widths of approximately 2 nm Centre: a bit-map image used as the input for a dip-pen nanolithography (DPN) routine. Right: AFM (lateral force) image of the resulting surface patterns
Schematic of dip-pen nanolithography, showing how the water meniscus is used to transport molecules to the surface.
Example of scanning thermal microscopy. Thermal conductivity image of a section from a glass filament/cyanate resin composite. The glass fibres clearly show greater thermal conductivity than the polymer matrix.

## Slide 29
AFM Basic steps
Preparation of samples
Place sample in instrument
Place probe in instrument
Optical aligment
Select initial settings
Select oscillation frequency
Probe approach
Feedback optimization
Scan and save images
Tip retract

## Slide 30
Blunt Probe
AFM Problems and Difficulties
Convex features such as particles tend to appear wider height may be accurate. Concave features such as pits tend to appear smaller (both less wide and less deep) with blunter probes.
Example of features appearing smaller due to the use of a blunt probe. Left: SEM image of a test pattern of squares. The sides of the squares are all equal. B: AFM image of the test pattern. Because the probe is not sharp, the test pattern squares appear much smaller than they should, and appear as rectangles instead of squares.
Illustration of the effect of using a blunt probe. These two images are of the same sample, and both are 1.5 μm x 1.5 μm x 40 nm. The image on the left was taken with a sharp probe, the image on the right with a blunt probe

## Slide 31
Broken and contaminated Probe tip (double tips)
AFM Problems and Difficulties
Left: an image of vesicles measured with a dirty tip. Centre: a badly broken and contaminated tip which produced double-tip images like these. Right: DNA molecules measured with a broken tip, each molecule has a false ‘twin’ next to it.

## Slide 32
Probe–sample angle problems
AFM Problems and Difficulties
Illustration of probe–sample angle problems. With the probe at an angle to the sample, distortions are introduced, and sample features appear asymmetric.
Effect of probe or probe side-wall imaging
Effect of probe or probe side-wall imaging. Top: illustrations of the effect of imaging a spike – an image of the probe is produced – and imaging a sphere-like feature – only at the top is the sample topography reproduced, and the rest of the image feature shows the probe’s side-walls.
Bottom: examples of probe–side-wall imaging. Left: red blood cells, right: S. aureus bacteria. In both cases, only the upper parts of the cells can be imaged correctly (some examples of probe side- wall imaging highlighted by arrows).

## Slide 33
Effect of x-y non-linearity in AFM images
AFM Problems and Difficulties
Effect of x-y non-linearity in AFM images. Left: example of an AFM image of a test sample when scanned without correctly linearizing the AFM scanner. Right: linearized AFM image of the same sample. The spacing of the squares at the top, bottom, left and right sides should be all the same distance apart.
Scanner Bow
Effect of scanner bow. Left: with tube scanners fixed at one end, the trajectory of the probe is curved. Right: the result is an apparent curvature in the height of measured samples, although the height change is small over a large area.
Graph showing the relationship between an actual z height and a measured z height in an AFM. Usually only one calibration point is measured as shown by the box, and the z piezoelectric is assumed to be linear, as shown by the dashed line. However, as is often the case, the piezoelectric actuator is non-linear, as shown by the solid line. In such cases incorrect z heights are measured unless the feature being measured has dimensions close to those of the calibration specimen.
one

## Slide 34
Edge overshoot
AFM Problems and Difficulties
Scanner creep
Scanner drift cause and effect. Left: creep in piezoelectric scanners causes the scanner to keep moving even after the applied voltage stops changing. Right: the effect on AFM images is most often seen as a distortion in the beginning of the scan (here, scanning from the top).
Edge overshoot in the z axis. Top: the probe is scanned from left to right across a feature on a surface; overshoot may be observed in the line profile at the leading and trailing edge of the features. Bottom: the AFM image of a test pattern appears to have no artefacts at first glance (left), but a line profile of the test pattern shows overshoot at the top of each of the lines (right, overshoot arrowed).

## Slide 35
Z angle measurement
AFM Problems and Difficulties
Other Problems
Illustration and example of errors in Z angle measurement by AFM caused by crosstalk. Top:
illustration of the effect. The sample has a series of repeating triangles at its surface. A line profile of the
sample shows that the triangles do not appear symmetric. Bottom: real AFM image of a sample having a
triangle pattern at its surface, and a line profile extracted from the AFM image. Although the angles of
the two facets are in reality equal, the AFM image suggests that this is not so.
During image processing
  Levelling
  Filtering
…noise
  Floor vibration
  Acoustic vibration
  Electric noise
  Vacuum leak
Feedback setting and scan rate
Surface contamination
Laser interference pattern
Sample drift

## Slide 36
Surface Roughness
AFM Applications
AFM height images of titanium surfaces subjected to various treatments. Samples were treated by A: polishing with colloidal silica, B: polishing followed by acid etching, C: sandblasting then treating with air plasma, and D: grinding with silicon carbide paper. All images have an X-Y scale of 20 x 20 μm, and a z scale of 6 μm.
Characterization of heterogeneous polymers
Examples of using phase imaging to characterize heterogeneous polymers. Left: phase imaging is sensitive to crystallinity via viscoelastic properties. In the height image (top left), it’s hard to distinguish the lamellae of the polymer due to macroscopic height differences. The phase image (below) shows them much more clearly. Right: example of phase imaging of a nanocomposite. Phase images of sliver nanoparticle/SBS nanocomposite films containing: (a) 0.5, (b) 3, (c) 7, and (d) 10 w.t% Ag nanoparticles.

## Slide 37
Nanodevice construction
AFM Applications
Example of electrical contacts assembled using the AFM as a manipulator. Left: carbon nanotube maneuvered to connect two gold electrodes.
Right: nanoparticles assembled into a nano- wire.
Electrical measurements of carbon nanotubes
Electrical measurements of carbon nanotubes. Top: charge injection into nanotubes. A: topography of isolated MWNT, and its EFM image before charging (B). After charge injection by the AFM probe (at the arrowed location), the EFM contrast reverses(C).

## Slide 38
Biology
AFM Applications
An example of real-time measurements of DNA-protein interactions. These images were obtained with high-speed IC-AFM, and the numbers in each frame indicate the time after imaging began. The images show a complex of EcoP15l, a DNA restriction enzyme, with DNA. The images from 1 to 10 seconds show the DNA loop passing through the enzyme.
Studies of bacterial morphology. Top left: Streptococcus, showing typical linear clusters. Top right: large clusters of Staphylococcus aureus. Bottom left: Salmonella biofilm showing pili-like fimbrial structures. Bottom right: E. coli.

## Slide 39
Commodity products application
AFM Applications
Examples of commercial products with applications in AFM. Left: 4 μm x 4 μm image of a razor blade. Middle: image of clay-containing coating on paper, showing hexagonal kaolin particles. Right: 35 μm x 35 μm image of human hair
Examples of applications in the data storage industry. Left: Three-dimensional representation of magnetic domains, in a 24 μm MFM image of a hard disk drive, allowing measurement of domain structure and spacing. Centre: 20 μm height image of a DVD. Automated software allows measurement of many properties from images such as this. Right: 16 μm height image of a laser bump in a hard disk landing zone. The scratch-like features are burnish marks, intentionally applied to allow air to escape while the head flies over the disk platter.
Micro-lens array, consisting of a surface on which a series of ca.7 μm lenses are manufactured. Dimensions, shape and roughness of these lenses are important parameters that the AFM can provide.